ABTEC-WI101-A0

A8产品技术规格书

| 产品型号 | ABTEC-A8 |
| --- | --- |
| 产品功能 | 音频测试分析仪 |
| 产品版本 | V1.0 |
| 产品外观 | 前面板图 |  |
|  | 后面板图 |  |

2020.04.29

| 目录TABLE OF CONTENTS |
| --- |
| 序号 NO. | 项目 ITEM | 页码 PAGE |
|  | 目录 | 1 |
| 1 | 适用范围 | 2 |
| 2 | 一般规范 | 2 |
| 3 | 功能模块 | 2 |
| 4 | 性能参数 | 3-10 |
| 5 | ACO选件 | 10 |
| 6 | 主观音质评测选件 | 11 |
| 7 | AX- BT 选件 | 12 |
| 8 | BT-DUO 选件 | 13 |
| 9 | AX- HDMI 选件 | 14 |
| 10 | AX- DSIO 选件 | 15 |
| 11 | 环境试验特性 | 16 |
| 12 | EMC特性 | 17 |
| 13 | ROHS 检测 | 18 |
| 14 | 产品遵守标准(国际标准/国家标准/行业标准/企业标准) | 18 |

一．适用范围

本产品适用于音频电路板及音频终端设备的测试应用。

二．一般规范

| 序号 | 项目参数 | 规格要求 |
| --- | --- | --- |
| 2.1 | 外形尺寸 | 428mm*430mm*120mm |
| 2.2 | 重量 | 8kg±0.5kg |
| 2.3 | 工作温度 | 0℃ ~ 45℃ |
| 2.4 | 工作电压（AC） | 220V,50Hz/110V220V,50H |
| 2.5 | 贮存温度 | -40ºC~75ºC |
| 2.6 | 湿度 | 10~80% |

- 功能模块

| 模拟输出通道数： | 8通道 |
| --- | --- |
| 模拟输入通道数： | 8通道 |
| 数字输出通道数： | 1通道 |
| 数字输入通道数： | 1通道 |
| 模拟输出端口类型： | D-SUB25/BNC |
| 模拟输入端口类型： | D-SUB25/BNC |
| 数字输出端口类型： | XLR/BNC/OPT |
| 数字输入端口类型： | XLR/BNC/OPT |
| 音视频输入端口类型: | AUX IN/OUT |
| 通讯端口类型： | USB/串口 |

四．性能参数

| Characteristic | Specifications | Supplemental Information |
| --- | --- | --- |
| ANALOG GENERATOR |
| Number of Channels | 8, independent amplitude control |  |
| Waveforms | Sine, sine split frequency, sine split phase, sine+DC offset, continuously swept-sine, noise, IMD signals, multi-tone, wave file playback |  |
| Sine Characteristics |
| Frequency Range (Fs) | 5 Hz to 80.1 kHz | Setting resolution is typically 45 µHz |
| Frequency Accuracy | ±(0.0003% + 100 μHz) |  |
| Amplitude Range | 0 to 14.4 Vrms [40.72 Vpp], bal; 0 to 7.20 Vrms [20.36 Vpp], unba |  |
| Amplitude Accuracy, 1 kHz |  |  |
| +15°C to +30°C | ±0.03 dB [±0.35%] |  |
| 0°C to +45°C | ±0.05 dB [±0.58%] |  |
| Flatness (1 kHz ref) |  |  |
| Fs = 10 Hz to 20 kHz | ±0.008 dB | typically <0.003 dB. |
| Fs = 20 kHz to 50 kHz | ±0.030 dB |  |
| Fs = 50 kHz to 80 kHz | ±0.10 dB |  |
| Residual THD+N(1,2) |  |  |
| Fs = 20 Hz–20 kHz | ≤(-103 dB + 1.4 µV) |  |
| Non-Harmonic Content |  | Typically <–110 dB when Fs ≤75 kHz, increasing to ≈ –55 dB at Fs =80 kHz |
| Phase Offset Range (SplitPhase) | –179.999 to +180.000 deg |  |
| DC Offset Range | ±12.00 Vdc balanced; ±6.00 Vdc unbalanced | DC offset limits maximum ac signal |
| Residual DC Offset | ≤0.25% of Vrms setting [≤0.09% of Vpp setting] + 100 μV |  |
| Noise Characteristics |
| Shape | White (<5 Hz to >80 kHz), Pink (<10 Hz to >80 kHz), IEC 60268-1 or BS EN 50332-1 |  |
| Amplitude Range | 0 to 40.72 Vpp, balanced; 0 to 20.36 Vpp, unbalanced | Amplitude calibration is approximate |
| IMD Test Signals (requires option AML) |
| SMPTE & MOD |
| Lower Frequency (LF) | 40 Hz to 1.00 kHz | LF tone must be ≤1/6 • HF tone. |
| Upper Frequency (HF) | 2.00 kHz to 60.00 kHz |  |
| MOD Upper Frequency (HF) | 240 Hz to 60.00 kHz |  |
| Mix Ratio (LF:HF) | 10:1, 4:1 or 1:1 |  |
| Amplitude Range | 0 to 40.72 Vpp, balanced; 0 to 20.36 Vpp, unbalanced. |  |
| Amplitude Accuracy | ±0.06 dB [±0.70%] |  |
| Residual IMD(1,2,3) | ≤0.0025% [-92 dB], 4:1 mix ratio |  |
| DFD |
| Difference Frequency (Fdiff) | 80 Hz to 2.00 kHz | Fmean = (F1 + F2)/2. |
| Mean Frequency (Fmean) | 250 Hz to 60.00 kHz | Fdiff = ¦F2 - F1¦ Fmean must be ≥6 • Fdiff |
| Amplitude Range | 0 to 40.72 Vpp, balanced; 0 to 20.36 Vpp, unbalanced. |  |
| Amplitude Accuracy | ±0.06 dB [±0.70%] |  |
| Residual IMD(1,2,3) | ≤0.0010% [-100 dB] |  |
| Multitone, Wave File Playback |
| Sample Rate Range (SR) | 8 kS/s to 108 kS/s, and 175 kS/s to 192 kS/s | Operation from 109 kS/s to 175 kS/s is possible, but with degraded flatness |
| Maximum File Size | 32 MSample. |  |
| Amplitude Range | 0 to 45.2 Vpp, balanced; 0 to 22.6 Vpp, unbalanced. | “.Wav” file must peak at digital full scale to obtain selected amplitude |
| Flatness (1 kHz ref) |  |  |
| SR = 175 kS/s to 192 kS/sec |  | Typically <0.012 dB to 20 kHz |
| SR = 8 kS/s to 108 kS/s |  | Typically <0.04 dB to 20 kHz; max frequency limited to ≈0.45 • SR |
| Spurious Content |  | Typically <-100 dB |
| Output Equalization | Arbitrary 30-pole output filter | The EQ operates on the first two internal generator channels, and is disabled for >2 output channels. |
| Source Resistance (Rs) |
| Balanced | 100Ω±1% | Grounded, symmetrical |
| Unbalanced | 50Ω±2%, | Electronically floating, 0.3 Vpk max; bnc shield to ground ≈10-17Ω ¦¦ 22nF |
| Max Output Current |  | Typically >30 mA peak, 10 mA dc; sum of all outputs ≤180 mA peak |
| Reverse Overload Protection |  | Up to 30 W |
| Output Related Crosstalk1 |  |  |
| Balanced | ≤(-100 dB + 1 µV) to 20 kHz |  |
| Unbalanced | ≤(-115 dB + 1 µV) to 20 kHz |  |
| ANALOG ANALYZER |
| Number of Channels | 8, independently auto-ranging. |  |
| Maximum Rated Input | 160Vpk, 120 Vdc, any input to ground; 0.5 Vpk bnc shields to ground |  |
| Input Impedance |
| Balanced | 100 kΩ‖≈230 pF, each side to ground |  |
| Unbalanced | 100 kΩ‖≈230 pF to bnc shield | Electronically floating, 0.5 Vpk max; bnc shield to ground ≈500Ω ¦¦ 22nF |
| Input Coupling | DC | Typically <0.5 μA bias current |
| Input Ranges | 320 mV to 100 V, 10 dB steps | Maximum ac signal ≈115 Vac, unbal or bal, in the 100 V range |
| Common Mode Rejection (4) |  | Max common mode signal range: |
| 320 mV, 1 V, 3.2 V ranges | ≥70 dB, 5 kHz to 20 kHz | ±6 Vpk |
| 10 V range | ≥50 dB, 5 Hz to 20 kHz | ±16 Vpk |
| 32 V range | ≥50 dB, 5 Hz to 20 kHz | ±60 Vpk |
| 100 V range | ≥45 dB, 5 Hz to 20 kHz | ±160 Vpk |
| Input Related Crosstalk |  | Typically <100 dB to 20 kHz between any two channels |
| Level (Amplitude) Measurement |
| Range | < 1 μV to 115 Vrms |  |
| Accuracy (1 kHz) |
| +15C to +30C | ±0.03 dB [±0.35%] |  |
| 0C to +45C | ±0.05 dB [±0.58%] |  |
| Flatness (1 kHz ref, DC coupling) |  |  |
| 10 Hz to 20 kHz | ±0.008 dB | Typically < 0.003 dB |
| 20 kHz to 50 kHz | ±0.030 dB |  |
| 50 kHz to 80 kHz | ±0.10 dB |  |
| Residual Noise (inputs shorted) | ≤1.3 μVrms, 20 kHz BW | Typically <8.0 nV / √Hz at 1 kHz. |
| THD+N Measurement |
| Fundamental Range | 5 Hz to 90 kHz |  |
| Measurement Range | 0 to 100% |  |
| Accuracy | ±0.5 dB |  |
| Residual THD+N(1,2) |  |  |
| 20 Hz–20 kHz fundamentals | ≤(-103 dB + 1.3 µV, 20 kHz BW); ≤ (-95 dB + 2.5 µV, 80 kHz BW); | Typically <-108 dB at 1 kHz, 2.5V |
| Bandwidth Limiting Filters |
| High-Pass(6) |  |  |
| DC | DC coupling |  |
| AC (< 10 Hz) | AC coupling | Response is 2-pole via a combination of analog and digital filters, and is typically -3 dB at 4.1 Hz |
| Butterworth | FHP (-3 dB) = 10 Hz to 90 kHz, 4-pole |  |
| Elliptic | FHP (-0.01 dB) = 10 Hz to 90 kHz; 5-pole; 0.01 dB pass-band ripple; ≤-60 dB stop-band |  |
| Low-Pass 5, 6 |
| ADC Passband | No filter is implemented, bandwidth and response are limited by the A/D and sample rate (SR) | -3 dB at ≈ 0.490 • SR, SR≤216 kS/s; |
| 20k (AES17), 40k (AES17) | Special filters conforming with AES17 |  |
| Butterworth | FLP (-3 dB) = 10 Hz to 90 kHz, 8-pole | ENBW ≈ 1.006 •FLP |
| Elliptic | FLP (-0.01 dB) = 10 Hz to 90 kHz, 8-pole;0.01 dB pass-band ripple; ≤-60 dB stop-band | ENBW ≈ 1.006 •FLP （varies due to warping） |
| Weighting | A-wt, B-wt, C-wt, CCIR-1k, CCIR-2k, CCITT, C-message, 50 μs or 75 μs de-emph (with and without A-wt), or None | Weighting filter is cascaded with both high-pass and low-pass filters |
| Input Equalization | Arbitrary 30-pole input filter | The EQ operates on any selected analyzer input channels. |
| IMD Measurement |
| Test Signal Compatibility |
| SMPTE & MOD | Any combination of 40 Hz–1 kHz (LF) and 240 Hz–60 kHz (HF), mixed in any ratio from 1:1 to 10:1 (LF:HF) | LF tone must be ≤1/6 • HF tone. |
| DFD | Any two-tone combination with mean frequency of 250 kHz–60 kHz and a difference frequency of 80 Hz–2.0 kHz | Fmean = (F1 + F2)/2. Fdiff = ¦F2 - F1¦ Fmean must be ≥6 • Fdiff |
| IMD Measured |  |  |
| SMPTE | Amplitude modulation of HF tone | Measurement BW is typ. 40-500 Hz. |
| MOD | d2, d3, d2+d3, or d2+d3+d4+d5 | Use“d2+d3”for measurements per IEC 60268 |
| DFD | d2, d3, d2+d3, or d2+d3+d4+d5 | Use“d2+d3”for measurements per IEC 60268 |
| CCIF | d2 only | “CCIF” is an archaic form of DFD that measures only the d2 product. CCIF uses a different 0 dB reference giving readings 2x higher than DFD. |
| Measurement Range | 0 to 20% |  |
| Accuracy | ±0.5dB |  |
| Residual IMD(1,2,3) |  |  |
| SMPTE & MOD | ≤-95 dB [0.0018%], 4:1 mix ratio |  |
| DFD | ≤-106 dB [0.0005%] |  |
| Frequency Measurement |
| Range | <5 Hz to 90 kHz, standard; |  |
| Accuracy | ±(0.0003% + 100 µHz) | Vin must be≥5 mV. |
| Resolution | 6 digits |  |
| Phase Measurement |
| Ranges | -90 to +270, ±180, or 0 to 360 deg |  |
| Accuracy | ±0.2 deg, 5 Hz to 5 kHz; ±0.8 deg, 5 kHz to 20 kHz; ±2.0 deg, 20 kHz to 50 kHz | Vin must be ≥5 mV, all channels |
| Resolution | 0.001 deg |  |
| DC Voltage Measurement | Valid only for input bandwidths ≤90k |
| Input Ranges | 0.32V to 100V, 10 dB steps | ±120 Vdc maximum in 100 V range |
| Accuracy |  |  |
| 0.32V range | ±(0.7% reading + 800 µV) |  |
| 1 V-100 V ranges | ±(0.7% reading + 0.1% range) |  |
| Normal Mode Rejection |  | Typically > 90 dB, 20 Hz to 20 kHz.. |
| DIGITAL I/O |
| DIGITAL OUTPUT RELATED: |
| Formats |  |  |
| Electrical, unbalanced | SPDIF-EIAJ per IEC60958, |  |
| Electrical, balanced | AES-EBU per AES3-1992, |  |
| Optical | Toslink® or equivalent |  |
| Sample Rate (SR) Range |  |  |
| Electrical | 27 kS/s to 200 kS/s | Usable over the extended range of 16 kS/s to 216 kS/s with degraded waveform |
| Optical | 27 kS/s to 108 kS/s |  |
| Sample Rate (SR) Accuracy | ±0.0003% [3 ppm] |  |
| Channel Status Bits | Full implementation per IEC-60958 automatically set, all channels same |  |
| User Bits and Validity Flag | Fully settable |  |
| Residual Jitter 1 |  |  |
| Electrical |  | Typically < 1.5 ns |
| Optical |  | Typically < 2.5 ns, SR ≤ 96 kS/s |
| EMBEDDED OUTPUT SIGNAL RELATED |

五．电声选件 AX- ACO

A8扬声器研发/生产测试选件，以单次快速扫描和易用性提供多参数检查扬声器的规格和性能，确保扬声器和音箱的完整性，并同时检查Rub & Buzz 缺陷、相对电平、相位、失真比和失真噪声电平等，比较黄金样本 。

通过快速连续扫描信号，找出高振幅因数和高峰值比从而检查 Rub & Buzz 异音缺陷。这些缺陷包括音圈移位、松散在音圈隙、不完全粘连的悬浮物料等；在这些情況下，所产生的问题是瞬变和很难检测到，不能使用标准计算例如 FFT测量或人工检测。

主要功能:

☆基本电声参数测量 （阻抗和 T/S小参数 ）

☆SPL/THD测试

☆喇叭生产测试

☆异音Rub & Buzz 检测

☆瀑布和极化图

☆空气泄露检测

☆其他声学测量

六. 主观音质评测选件

AX-PESQ     基于ITU-T P.862标准

AX-POLQA2   基于 ITU-T P.863标准

感知音频测量广泛应用与智能手机，VoIP 网络和设备测试，衡量人们如何感知声音品质。感知声音品质测试是特別有用于宽频或位元速率有限的设备，使用标准的正弦波测试如频响和失真等这些参数不能有效地测试和比较不同设备性能；最后，只有用感知和电气音频测试組合，才可以完全説明這些设备的性能。

PESQ是一个感知语音品质测量工具。它由德国OPTICOM 有限公司开发，根据 ITU-T 建议专为低频宽的设备，例如手机和智慧手机上测试语音品质，适当的测试信号可以达到类似使用人类测试的结果，有很高的相关性；

POLQA2 加強至宽带语音品质和声学接口等等测试。

- AX- BT 选件

| 5.1.1 | 蓝牙协议版本 | 2.1+EDR |
| --- | --- | --- |
| 5.1.2 | Profiles/Roles Supported | A2DP Source |
|  |  | A2DP Sink |
|  |  | HFP Audio Gateway |
|  |  | HFP Hands-Free |
|  |  | HSP Audio Gateway |
|  |  | HSP Headset |
|  |  | AVRCP Controller |
| 5.1.3 | Codecs Supported | SBC |
|  |  | APT-X |
| 5.1.4 | RF端口类型 | Type N, Female |
|  |  | BT Antenna (2.4GHz) |
| 5.1.5 | RF 输入阻抗 | 50 Ω |
| 5.1.6 | RF 输出阻抗 | 50 Ω |
| 5.1.7 | RF 功率 | 0 dBm;最大可以达到 +4 dBm |
| 5.1.8 | RF 灵敏度(0.1% BER) | –81 dBm |

- BT-DUO

BT-DUO是一个经过完全修订、升级的蓝牙硬件模块，具有专用的源和接收器、新的蓝牙芯片、最高固件、更高的RF功率和改进的RF屏蔽。并且提供了一系列新的音频编解码器，支持配置文件中的新操作功能，以及更快的连接时间。使工程师能够对其蓝牙设备进行直接测量，消除蓝牙Dongle的不确定性和不便性，使蓝牙音频测试更快速、更简单、更可靠。

主要功能：

蓝牙设备核心版本：V4.2

配置文件版本：A2DP v1.3，AVRCP v1.4，HFP v1.7，HSP v1.2

A2DP音频编解码器：SBC，aptX，aptX-LL，aptX-HD，AAC，LHDC(HWA)

HFP音频编解码器：CVSD，mSBC

RF连接：SMA(母座）

射频输入阻抗：典型值50 Ω

射频输出阻抗：典型值50 Ω

射频功率：典型最大值+8 dBm

射频灵敏度（0.1% BER）：Min  -90 dBm

九．AX- HDMI

AX- HDMI选件（HDMI+音频回传通道（ARC）是一种多功能的音频监视和音频测量的功能选件,它功能强大，使用简单方便,测量结果准确，可广泛应用于多媒体产品的音频质量测试，可提供的所有标准音频测量值包括电平、信噪比（SNR）、失真、相位、串音、群延迟等。

主要功能：

执行标准：HDMI1.3；

接口类型：type A；

通道数：2，8

比特：8bit 到24bit

采样率：可达192kHz

支持格式：PCM、Dolby Digital、DTS等

具有对可支持接收器和电视的ARC音频质量和可连接性的功能；

具有生成线性/PCM音频流，支持从编码前的音频测试文件中生成无损格式

(比如杜比TrueHD™和dts-HD主音频™)和压缩格式（比如杜比数字™和dts数字环绕立体声™)；

具有兼容性和减采样/缩混/转码功能；

支持高清晰度多媒体接口音频回传信道；

具有查看查看和编辑HDMI 增强扩展显示识别数据(E-EDID)；

支持多个EDID配置；

具有生成视频信号（分辨率达到1080P)和第三方视频支持。

十．AX-DSIO

AX-DSIO 数字串行输入/输出选件是 新一代多通道芯片级音频接口，可广泛应用到各种设备和工程测试中。

关键参数指标：

脉冲电压：1.8V;2.5V;3.3V              格式：左对齐；右对齐；I2S;DSP

数据长度：8—32位                   抖动：选择ON或OFF

字长：    8—128位                  主时钟源：发射器（外部或内部);接收器（外部或内部)

采样率：4KHz—216KHz               比特/时钟方向：发射器（IN或OUT); 接收器（IN或OUT)

主时钟速率：4KHz—56MHz            主时钟：可反转

位时钟边缘同步：上升或下降

多通道配置：TDM (1，2，4，6，8 个通道）；多个数据连接（1，2，4，6，8 通道）

十一. 环境试验特性

环境试验特性

| 试验条件 | 环境温度 | 15℃～35℃ |
| --- | --- | --- |
|  | 相对湿度 | 25%～75% |
|  | 气压 | 86~106kpa |
| 7.1 | 电源适应性 | 执行 GB/T 6587-2012 条款4.10 |
| 7.2 | 介电常数 | 执行 GB/T 6587-2012条款5.8.2 |
| 7.3 | 保护接地 | 执行 GB/T 6587-2012条款5.8.3 |
| 7.4 | 温度实验 | 执行 GB/T 6587-2012条款5.9.1 |
| 7.5 | 湿度实验 | 执行 GB/T 6587-2012条款5.9.2 |
| 7.6 | 冲击实验 | 执行 GB/T 6587-2012条款5.9.4 |
| 7.7 | 裸机振动试验 | 执行 GB/T 6587-2012条款5.9.3 |
| 7.8 | 包装振动实验 | 执行 GB/T 6587-2012条款5.10.1 |

十二．EMC 特性

执行 GB/T 6587-2012 《电子测量仪器通用规范》---4.9条款

| 8.1 | 传导骚扰 | 150kHz～30MHz电源端子的传导骚扰电压限值（B级） |
| --- | --- | --- |
| 8.2 | 辐射骚扰 | 辐射骚扰限值 B级 |
| 8.3 | 静电放电抗扰度 | 性能判据：等级B |
| 8.4 | 射频电磁场辐射抗扰度 | 性能判据：等级A |
| 8.5 | 电快速瞬变脉冲群抗扰度 | 实验判据：等级B (GB/T 17626.4-2008 等级四) |
| 8.6 | 浪涌（冲击）抗扰度 | 实验判据：等级B |
| 8.7 | 射频场感应的传导骚扰抗扰度 | 性能判据：等级A |
| 8.8 | 电压跌落、暂降和短时中断抗扰度 | 性能判据：等级B |

十三．ROHS 检测

本产品符合GB/T26572-2011对有害物质的限量要求

备注：

O:表示该有害物质在该部件所有均质材料中的含量均在GB/T26572-2011规定的限量要求以下。

×：表示该有害物质至少在该部件的某一均质材料中的含量超出GB/T26572-2011规定的限量要求。

十四．产品遵守标准(国际标准/行业标准/企业标准)

本产品遵守如下国家标准：

10.1《GB_T 6587-2012_5625电子测量仪器通用规范》

10.2《JJF 1395-2013 音频分析仪校准规范》

10.3 《GB/T26572-2011电子电气产品中限用物质的限量要求》

10.4《SJ/T11364-2014 电子电气产品中限用物质限制使用标识要求》